## Slide 1: RRAM Array Bring-Up: TSMC 64K Cell (256x256)
Stanford University
2/19/21
Akash Levy
Notes: This presentation will attempt to kill two birds with one stone
It will be structured as a 15 min crash course on how to test an RRAM array, with the TSMC RRAM array as a case study
Assumptions: direct access to the memory array (with address decoders)

## Slide 2: READ
Step One

## Slide 3: READ Operation
READ requires source and measure
Need to source READ voltage on the BL
Need to measure READ current on the BL
R = VBL/IBL
Some systems will allow you to source and measure with same unit, making life easy
    Need to specify current limit correctly!
Otherwise, using shunt resistor and voltage measurement unit works well for high resistance
Load = 1T1R structure, shunt = constant resistor
    Want small value so IR drop is small
    Too small gives you low resolution signal
    Rshunt = 560Ω was reasonable choice
      Rload ∈ {7kΩ, 500kΩ} and voltage resolution of 0.1 mV
  (up to 1.1MΩ accurate)
Measure voltage here
Shunt resistor setup
VBL
Vshunt
First test with constant resistor as load!

## Slide 4: READ Waveform: Always Verify with Scope
Set VSL=0V
Set VWL=2.5V
    Higher VWL means less transistor resistance included in measurement
Set VBL=200mV
    Yields good signal with low read disturb
Find conservative settling times
    Address settling time (500us)
    VWL ON time (SW timed)
    VBL settling time (2ms)
    Sampling (100Sa @ 1 MSa/s)
NOTE: there are other valid waveforms that work
https://tinyurl.com/rram-read-waveform

## Slide 5: Single-Bit Operation Conditions
Operation | READ | FORM | SET | RESET
BL/SL Start (V) | 0.2 (BL) |  |  | 
BL/SL Stop (V) | N/A |  |  | 
BL/SL Step (V) | N/A |  |  | 
WL Start (V) | 2.5 |  |  | 
WL Stop (V) | N/A |  |  | 
WL Step (V) | N/A |  |  | 
WL Pulse Width (us) | 2000 |  |  | 
Stop Condition | N/A |  |  | 
Loop Order | N/A |  |  | 

## Slide 6: FORM
Step Two

## Slide 7: FORM Waveform: Always Verify with Scope
Before FORMing, it is not possible to test your READ waveform
The first step is to just apply a long SET pulse (will show in a few slides)
    Example: VBL=3V, VWL=4V, VSL=0V
Then READ, and check that resistance dropped significantly
    If it does, this is the first sign that your device is alive! 
Now try to do “dynamic FORM”:
    FORM_THRESH, FORM_PW = 50E3, 1000 # 50kOhm FORMing threshold, 1ms
    for VWL in range(0, 2, step=0.05): # VWL outer loop
    for VBL in range(2, 4, step=1): # VBL inner loop, optional (can fix VBL val)
    PULSE(VBL, VWL, VSL=0, FORM_PW) # FORMing pulse
    if READ() <= FORM_THRESH:
    print(‘FORM succeeded’)
    break
    else: # if loop completes without break
    print(‘FORM failed’)

## Slide 8: Full-Array FORMing Yield
Initial:
39 Failures…
Small number of FORM failures due to 1.5V VWL limit & 200us PW
Failures happen further down the BL/WL, probably due to IR drop
Important not to over-SET
    Do not increase VWL limit!
    Instead: increase PW or do rerun
Rerunning FORM for failed cells successfully FORMed 38/39
Rerunning a second time on the final failed cell FORMed it
Result: 100% FORM yield!
Afterwards: perhaps adjust PW

## Slide 9: Full-Array FORMing Statistics
Certain BLs seem to have higher VWL (every 8 or so) and hence lower resistance
Not sure why this might be the case… any ideas?

## Slide 10: Full-Array FORMing Statistics

## Slide 11: Single-Bit Operation Conditions
Operation | READ | FORM | SET | RESET
BL/SL Start (V) | 0.2 (BL) | 2 |  | 
BL/SL Stop (V) | N/A | 4 |  | 
BL/SL Step (V) | N/A | 0.5 |  | 
WL Start (V) | 2.5 | 0.9 |  | 
WL Stop (V) | N/A | 1.5 |  | 
WL Step (V) | N/A | 0.05 |  | 
WL Pulse Width (us) | 2000 | 1000 |  | 
Stop Condition | N/A | < 50kΩ |  | 
Loop Order | N/A | WL outer/ BL inner |  | 

## Slide 12: Single-Bit Program: SET/RESET
Step Three

## Slide 13: SET/RESET Waveform: Always Verify with Scope
Reuse dynamic FORM template
Pulse VWL
Find good timing
    Address settling time (500us)
    VBL/VSL ON time (SW timed)
    PW (10us)
    VWL OFF time (SW timed)
NOTE: there are other valid waveforms that work
https://tinyurl.com/rram-set-reset-waveform

## Slide 14: Dynamic SET/RESET & TARGET (Write-Verify)
  def dynamic_set(SET_THRESH=11E3, SET_PW=10):
    for VWL in range(0, 2, step=0.05):
    for VBL in range(2, 4, step=1):
    PULSE(VBL, VWL, VSL=0, SET_PW)
    R = READ()
    if R <= SET_THRESH:
    print(‘SET succeeded’)
    break
    else: print(‘SET failed’)
    return R
  def dynamic_reset(RST_THRESH=100E3, RST_PW=10):
    for VWL in range(0, 2, step=0.05):
    for VSL in range(2, 4, step=1):
    PULSE(VWL, VSL, VBL=0, RST_PW)
    R = READ()
    if R >= RST_THRESH:
    print(‘RESET succeeded’)
    break
    else: print(‘RESET failed’)
    return R
  def target(Rmin, Rmax, MAXI=5):
  for i in range(MAXI):
  R = READ()
  if R > Rmax:
  R = dynamic_set(Rmax)
  if R < Rmin:
  R = dynamic_reset(Rmin)
  if Rmin < R < Rmax:
  print(‘TARGET succeeded’)
  break
  else:
  print(‘TARGET failed’)
  return R

## Slide 15: Checkerboard: Random Access Pattern Check
Error with single cell FORMed at VWL=1.7V
Cell cannot be fully RESET, only up to 20kOhm

## Slide 16: “Anti-Checkerboard”: Random Access Pattern Check
3 errors when trying to flip
Cells cannot be fully RESET on first pulse, only up to 20-40 kOhm
Happens if cells are over-FORMed: shows importance of proper FORMing

## Slide 17: Operation conditions
Operation | READ | FORM | SET | RESET
BL/SL Start (V) | 0.2 (BL) | 2 | 2 | 3
BL/SL Stop (V) | N/A | 4 | 2 | 3
BL/SL Step (V) | N/A | 0.5 | N/A | N/A
WL Start (V) | 2.5 | 0.9 | 0.9 | 0.9
WL Stop (V) | N/A | 1.5 | 1.5 | 1.5
WL Step (V) | N/A | 0.05 | 0.1 | 0.1
WL Pulse Width (us) | 2000 | 2000 | 10 | 10
Stop Condition | N/A | < 50kΩ | Below target (11kΩ) | Above target (100kΩ)
Loop Order | N/A | WL outer/ BL inner | WL only | WL only
Total FORM Time:
~3 cells/s
=
7 hrs/array
Total Prog Time:
~4 cells/s
=
6 hrs/array

## Slide 18: Full-Array Programming (Iteration 1)

## Slide 19: Full-Array Programming (Iteration 2)

## Slide 20: Full-Array Programming (Iteration 3)

## Slide 21: Multi-Bit Programming
Step Four

## Slide 22: Find the SET/RESET “Limits”
For multi-bit programming, you care a lot about the dynamic range
    Indicates multi-bit tunability
Over-SET: if resistance goes too low, transistor will not be able to pass enough current to RESET
Device Shorting: happens if current is too high through the device
    Caused by VWL being too high
Find the minimum resistance and maximum VWL
Note that all these values may vary significantly die-to-die!

## Slide 23: Characterization Sweeps
ISPP Characterization
    Ramp VWL with fixed VBL/VSL starting from “fully SET/RESET state” and observe the trajectory
Other ramping sweeps
    VBL/VSL ramping with fixed VWL
Fixed-Pulse SET Characterization
    Try fixed voltage pulses instead of ramping
    This is useful for determining a good “starting point” for ISPP

## Slide 24: VWL SET Sweeps
Same as standard ISPP Algorithm

## Slide 25: ISPP: VWL SET Sweeps (Resistance, Large PW)

## Slide 26: ISPP: VWL SET Sweeps (Resistance, Small PW)

## Slide 27: ISPP: VWL SET Sweeps (Conductance, Large PW)
No READ disturb below 0.4V
Dynamic range seems to be ~0-120uS

## Slide 28: ISPP: VWL SET Sweeps (Conductance, Small PW)
Smaller PW gives more linear behavior
But more variability(?)

## Slide 29: VBL SET Sweeps

## Slide 30: VBL SET Sweeps (Resistance, Large PW)
SET around 1V
Fine Tuning

## Slide 31: VBL SET Sweeps (Resistance, Small PW)

## Slide 32: VBL SET Sweeps (Conductance, Large PW)

## Slide 33: VBL SET Sweeps (Conductance, Small PW)

## Slide 34: VWL RESET Sweeps

## Slide 35: VWL RESET Sweeps (Resistance, Large PW)
SHORT!
Highly abrupt and stochastic behavior

## Slide 36: VWL RESET Sweeps (Resistance, Small PW)
Do not push past VWL=3.5V, VSL=4V!
SHORT!
Highly abrupt and stochastic behavior
No response

## Slide 37: VWL RESET Sweeps (Conductance, Large PW)
Highly abrupt and stochastic behavior
SHORT!
“Safe” tuning window
No response
150 uS is too high, can’t be RESET
120 uS seems to be okay

## Slide 38: VWL RESET Sweeps (Conductance, Small PW)
No response

## Slide 39: VSL RESET Sweeps

## Slide 40: VSL RESET Sweeps (Resistance, Large PW)
Best Gradual RESET

## Slide 41: VSL RESET Sweeps (Resistance, Small PW)
No response
Abrupt RESET

## Slide 42: VSL RESET Sweeps (Conductance, Large PW)
Best Gradual RESET
Abrupt RESET

## Slide 43: VSL RESET Sweeps (Conductance, Small PW)

## Slide 44: Conductance Tuning Curves

## Slide 45: Conductance Tuning Curve (using TARGET)
Now we know the dynamic range of the RRAM
We also know good SET/RESET voltages to use
Using TARGET, we can now tune within our conductance window
Choose 0-128 uS and have 16 evenly spaced target ranges
Programming once will give bad relaxation
Do tail distribution reprogramming (TDR) to fix cells that quickly relax out of target range

## Slide 46: Conductance Tuning Curve (after relaxation)
Now we know the dynamic range of the RRAM
We also know good SET/RESET voltages to use
Using TARGET, we can now tune within our conductance window
Choose 0-128 uS and have 16 evenly spaced target ranges
Programming once will give bad relaxation
Do tail distribution reprogramming (TDR) to fix cells that quickly relax out of target range

## Slide 47: Conductance Tuning Curve (5 iters of TDR w/relaxation)
Now we know the dynamic range of the RRAM
We also know good SET/RESET voltages to use
Using TARGET, we can now tune within our conductance window
Choose 0-128 uS and have 16 evenly spaced target ranges
Programming once will give bad relaxation
Do tail distribution reprogramming (TDR) to fix cells that quickly relax out of target range
Even spacing of the ranges is probably the wrong thing to do
Need something like sigma-based allocation
Also better pulsing/pre-cycling…

## Slide 48: Conductance Tuning Curve: Confusion Matrix

## Slide 49: Short-Term Retention Behavior
Do many READs after programming and observe statistical behavior

## Slide 50: Retention using ISPP SET

## Slide 51: Retention using ISPP RESET

## Slide 52: Retention using TARGET

## Slide 53: Endurance Measurement
In progress
Done up to 100k cycles
Resistance window

## Slide 54: TODO
There’s always more…

## Slide 55: More Measurements/Experiments
Sigma-Based Allocation for better multi-bit programming
More detailed retention study…
Variability
    Cycle-to-cycle variations
      Cycle the same cell and view distribution
    Array location variations
      We have already investigated this for FORMing but not for other ops
      One challenge with the TCAM chip is that there are 4 different drivers for WLs and 2 different drivers for BLs which introduce further variation in the behavior
    Die-to-die variations
Non-linear Behavior
    Use different READ voltages and observe current
High-Temperature Retention
    Bake chip and observe changes in resistance
    Can extract activation energies and estimate device lifetime via extrapolation
MORE IDEAS?

## Slide 56: Thank you!
(If you are interested in the RRAM data or using the LabVIEW program, contact Akash at akashl@Stanford.edu)

## Slide 57: Me: Mom, can I have multi-bit RRAM? Mom: No, we have multi-bit RRAM at home Multi-Bit RRAM at Home: