## Slide 1: RRAM Array Testing Crash Course: TSMC 64K Cell Array (256x256)
Stanford University
2/19/21
Akash Levy
Notes: This presentation will attempt to kill two birds with one stone
It will be structured as a 15 min crash course on how to test an RRAM array, with the TSMC RRAM array as a case study
Assumptions: direct access to the memory array (with address decoders)

## Slide 2: READ
Step One

## Slide 3: READ Operation
READ requires source and measure
Need to source READ voltage on the BL
Need to measure READ current on the BL
R = VBL/IBL
Some systems will allow you to source and measure with same unit, making life easy
    Need to specify current limit correctly!
Otherwise, using shunt resistor and voltage measurement unit works well for high resistance
Load = 1T1R structure, shunt = constant resistor
    Want small value so IR drop is small
    Too small gives you low resolution signal
    Rshunt = 560Ω was reasonable choice
      Rload ∈ {7kΩ, 500kΩ} and voltage resolution of 0.1 mV
  (up to 1.1MΩ accurate)
Measure voltage here
Shunt resistor setup
VBL
Vshunt
First test with constant resistor as load!

## Slide 4: READ Waveform: Always Verify with Scope
Set VSL=0V
Set VWL=2.5V
    Higher VWL means less transistor resistance included in measurement
Set VBL=200mV
    Yields good signal with low read disturb
Find conservative settling times
    Address settling time (500us)
    VWL ON time (SW timed)
    VBL settling time (2ms)
    Sampling (100Sa @ 1 MSa/s)
NOTE: there are other valid waveforms that work
https://tinyurl.com/rram-read-waveform

## Slide 5: Single-Bit Operation Conditions
Operation | READ | FORM | SET | RESET
BL/SL Start (V) | 0.2 (BL) |  |  | 
BL/SL Stop (V) | N/A |  |  | 
BL/SL Step (V) | N/A |  |  | 
WL Start (V) | 2.5 |  |  | 
WL Stop (V) | N/A |  |  | 
WL Step (V) | N/A |  |  | 
WL Pulse Width (us) | 2000 |  |  | 
Stop Condition | N/A |  |  | 
Loop Order | N/A |  |  | 

## Slide 6: FORM
Step Two

## Slide 7: FORM Waveform: Always Verify with Scope
Before FORMing, it is not possible to test your READ waveform
The first step is to just apply a long SET pulse
    Example: VBL=3V, VWL=4V, VSL=0V
Then READ, and check that resistance dropped significantly
    If it does, this is the first sign that your device is alive! 
Now try to do “dynamic FORM”:
    FORM_THRESH, FORM_PW = 50E3, 1000 # 50kOhm FORMing threshold, 1ms
    for VWL in range(0, 2, step=0.05): # VWL outer loop
    for VBL in range(2, 4, step=1): # VBL inner loop, optional (can fix VBL val)
    PULSE(VBL, VWL, VSL=0, FORM_PW) # FORMing pulse
    if READ() <= FORM_THRESH:
    print(‘FORM succeeded’)
    break
    else: # if loop completes without break
    print(‘FORM failed’)

## Slide 8: Full-Array FORMing Yield
Initial:
39 Failures…
Small number of FORM failures due to 1.5V VWL limit & 200us PW
Failures happen further down the BL/WL, probably due to IR drop
Important not to over-SET
    Do not increase VWL limit!
    Instead: increase PW or do rerun
Rerunning FORM for failed cells successfully FORMed 38/39
Rerunning a second time on the final failed cell FORMed it
Result: 100% FORM yield!
Afterwards: perhaps adjust PW

## Slide 9: Full-Array FORMing Statistics
Certain BLs seem to have higher VWL (every 8 or so) and hence lower resistance
Not sure why this might be the case… any ideas?

## Slide 10: Full-Array FORMing Statistics

## Slide 11: Operation Conditions
Operation | READ | FORM |  | 
BL/SL Start (V) | 0.2 (BL) | 2 |  | 
BL/SL Stop (V) | N/A | 4 |  | 
BL/SL Step (V) | N/A | 0.5 |  | 
WL Start (V) | 2.5 | 0.9 |  | 
WL Stop (V) | N/A | 1.5 |  | 
WL Step (V) | N/A | 0.05 |  | 
WL Pulse Width (us) | 2000 | 1000 |  | 
Stop Condition | N/A | < 50kΩ |  | 
Loop Order | N/A | WL outer/ BL inner |  | 

## Slide 12: Single-Bit Program: SET/RESET
Step Three

## Slide 13: SET/RESET Waveform: Always Verify with Scope
Reuse dynamic FORM template
Pulse VWL
Find good timing
    Address settling time (500us)
    VBL/VSL ON time (SW timed)
    PW (10us)
    VWL OFF time (2ms)
NOTE: there are other valid waveforms that work
https://tinyurl.com/rram-set-reset-waveform

## Slide 14: Operation conditions
Operation | READ | FORM | SET | RESET
BL/SL Start (V) | 0.2 (BL) | 2 | 2 | 3
BL/SL Stop (V) | N/A | 4 | 2 | 3
BL/SL Step (V) | N/A | 0.5 | N/A | N/A
WL Start (V) | 2.5 | 0.9 | 0.9 | 0.9
WL Stop (V) | N/A | 1.5 | 1.5 | 1.5
WL Step (V) | N/A | 0.05 | 0.1 | 0.1
WL Pulse Width (us) | 2000 | 2000 | 10 | 10
Stop Condition | N/A | < 50kΩ | Below target (9-11kΩ) | Above target (70-90kΩ)
Loop Order | N/A | WL outer/ BL inner | WL only | WL only
Total FORM Time:
3 cells/s
=
7 hrs/array
Total Prog Time:
4 cells/s
=
6 hrs/array

## Slide 15
  def dynamic_set():
    SET_THRESH, SET_PW = 11E3, 10
    for VWL in range(0, 2, step=0.05):
    for VBL in range(2, 4, step=1):
    PULSE(VBL, VWL, VSL=0, SET_PW)
    if READ() <= SET_THRESH:
    print(‘SET succeeded’)
    break
    else: print(‘SET failed’)

## Slide 16: Checkerboard: Random Access Pattern Check
Error with single cell FORMed at VWL=1.7V
Cell cannot be fully RESET, only up to 20kOhm

## Slide 17: “Anti-Checkerboard”: Random Access Pattern Check
3 errors when trying to flip
Cells cannot be fully RESET on first pulse, only up to 20-40 kOhm
Happens if cells are over-FORMed: shows importance of proper FORMing

## Slide 18: Full-Array FORMing Yield
Initial:
39 Failures…
Small number of FORM failures due to 1.5V VWL limit
Failures happen further down the BL/WL, probably due to IR drop
Important not to over-SET
    Cannot increase limit!
Rerunning FORM for failed cells successfully FORMed 38/39
Rerunning a second time on the final failed cell FORMed it
Result: 100% FORM yield!
Afterwards: realize longer PW will always give 100% yield

## Slide 19: Full-Array FORMing Statistics
Certain BLs seem to have higher VWL and hence lower resistance

## Slide 20: Full-Array FORMing Statistics

## Slide 21: Status
Done with FORM, SET, RESET, Target Resistance
Done with checkerboard tests
Done with bitstream programming
Speed + yield optimizations
Done with full array FORM
Done with full array yield analysis
Done with full array programming of one array (first iteration)
In progress:
    Tail distribution reprogramming
    Process more arrays with different LRS/HRS
    Finalize figures

## Slide 22: Operation conditions
Operation | READ | FORM | SET | RESET
BL/SL Start (V) | 0.2 (BL) | 2 | 2 | 3
BL/SL Stop (V) | N/A | 4 | 2 | 3
BL/SL Step (V) | N/A | 0.5 | N/A | N/A
WL Start (V) | 2.5 | 0.9 | 0.9 | 0.9
WL Stop (V) | N/A | 1.5 | 1.5 | 1.5
WL Step (V) | N/A | 0.05 | 0.1 | 0.1
WL Pulse Width (us) | 2000 | 2000 | 10 | 10
Stop Condition | N/A | < 50kΩ | Below target (9-11kΩ) | Above target (70-90kΩ)
Loop Order | N/A | WL outer/ BL inner | WL only | WL only
Total FORM Time:
3 cells/s
=
7 hrs/array
Total Prog Time:
4 cells/s
=
6 hrs/array
Switched to analog driver

## Slide 23: Multi-Bit Programming
Step Four